## Slide 1: Principles of Mechanical Ventilation
Matt Hegewald, MD
Intermountain Medical Center
University of Utah

## Slide 2: Principles of Mechanical Ventilation
Design features
  Breath design
  Modes
Respiratory system mechanics

## Slide 3: Spontaneous vs. Positive-Pressure Breathing
Pressure
Positive Pressure
Volume
Pressure
Time
Spontaneous
I
E
I
E
Notes: 15
Let’s examine the pressure and volume curves typical of a spontaneous and a positive-pressure breath delivered via a mechanical ventilator. On the left, which is the spontaneous breath, we can see the application of information we just discussed. The inspiratory phase has a slight negative pressure; then, on exhalation there is a positive deflection. From our recent discussion of positive-pressure ventilation, notice that the pressure remains positive during the entire inspiratory phase. It is easy to see how spontaneous ventilation, which generates a negative intrathoracic pressure, aids venous blood flow to the heart. Also, positive-pressure ventilation may decrease venous blood flow, and thus some patients may require intravenous volume.
Now, we'll take a moment to define ventilator parameters and then review some of the basic ventilator settings.

## Slide 4: Breath Characteristics 3 “Ts”
Trigger
  What initiates a breath
  Machine timer (control) vs patient effort (assist)
Gas delivery target or limit
  What governs gas flow
  Flow target vs pressure target
Cycle
  What terminates the breath
  Volume, time, flow, pressure

## Slide 5: Breath Characteristics Trigger
Controlled – machine timer
Assisted – patient effort
  Pressure trigger – effort produces pressure drop in vent circuit
  Flow trigger – effort draws gas out of a continuous flow through the vent circuit

## Slide 6
L/sec
L
cm/H2O
cm/H2O
Flow
Vol
Paw
Pes
A
B

## Slide 7: Breath Characteristics Gas Delivery Target
Airway
Pressure
Flow
Volume
0
+
-
0
in
out
increasing
Set pressure
Either Set ti or
Set minimal flow
(cycle off)
Set flow
Set volume
(cycle off)

## Slide 8: Understanding the Flow-Time Waveform
There are two components to the flow-time waveform
  The inspiratory arm:
    Active in nature
    The character is affected by the ventilator mode, patient effort, and total system resistance and compliance.
  The expiratory arm:
    Passive in nature
    The character is determined mainly by elastic recoil of the patients lungs and airway resistance.
    Also affected by patient respiratory effort (if any)
There are two commonly used types of flow patterns available on most ventilators
  The ’square wave’ or ‘constant flow’ pattern
  The ‘ramp’ (decelerating) type pattern

## Slide 9: The‘Square Wave’Flow Pattern
Time
Flow
Inspiratory
arm
Expiratory
arm
Inspiratory time = Tidal volume
Flow rate
The expiratory flow is passive and is
determined by airways resistance and the elastic recoil of the lungs
The inspiratory flow rate remains constant over the entire inspiration

## Slide 10: The‘Decelerating Ramp’Flow Pattern
Time
Flow
Inspiratory
arm
Expiratory
arm
Inspiratory time = Tidal volume
Flow rate
For a given tidal volume, the inspiratory time is higher in this type of flow pattern as compared to the square wave pattern
The inspiratory flow rate decelerates as a function of time to reach zero flow at end inspiration

## Slide 11
1
2
3
4
5
6
30
Paw
cmH2O
Sec
V
.
Trigger Effort
PIP
Ti
-10
Notes: 12
12
12
12
12
12
The Pressure-Time curve depicts not only the shape of the curve but displays any PEEP present, the PIP, as well as any inspiratory trigger effort as seen in the second pressure curve.
This Pressure-Time curve shows the different types of pressure curves associated with different types of flow delivery.
a. the linear rapid rise to pressure produced by a square
flow wave
b. the curved gradual rise to pressure resulting from a
descending flow wave
c. the curved “shark’s fin” pressure wave with a long
inspiratory plateau

## Slide 12: Breath Characteristics Cycle
Three common types
  Reach set volume
  Reach set time
  Reach specified flow reduction
Airway pressure as “backup” cycle
  Pressure alarm and cycle with occlusion or expiratory efforts

## Slide 13: Breath Characteristics
 | TRIGGER | TARGET/LIMIT | CYCLE
Volume control (VC) Volume assist (VA) Pressure control (PC) Pressure assist (PA) Pressure support (PS) | Time Effort Time Effort Effort | Flow Flow Pi** Pi** Pi** | Vol* Vol* Time* Time* Flow*
Airway
Pressure
Flow
Volume
0
+
-
0
in
out
increasing
Set pressure
Set flow
Set volume
Set flow
Set volume
Set Ti
Set pressure
Set Ti
Set pressure
Set minimal
flow
Machine
triggered
Machine
triggered
Patient
triggered
Patient
triggered
Patient
triggered

## Slide 14: Ventilator Modes
Volume Assist/Control (VAC)
  Set parameters: inspiratory flow rate (usually 60 L/min), frequency, Vt, PEEP, FIO2
  Flow, Vt and frequency determine TI, TE and I:E ratio
  Patient triggered breaths result in same flow and Vt as control breaths – this changes I:E ratio and can create auto-PEEP
  Advantages: constant Vt guaranteed, mode used in ARMA trial demonstrating reduced mortality in ARDS
  Pplat = Vt/Crs + PEEP

## Slide 15: Ventilator Modes
Pressure Assist/Control (PCV)
  Set parameters: Inspiratory pressure, I:E ratio, frequency, PEEP, FIO2
  Inspiratory flow rate is matched to patient effort – may improve patient-ventilator synchrony and comfort
  Variable Vt (lung-protective Vt not assured) but predictable Ppeak and Pplat (could lessen VILI); Ppeak may not equal alveolar pressure with obstruction
  Vt = (Pinsp-PEEP) x Crs

## Slide 16: Positive Pressure Ventilation VAC vs. PCV
VAC
Volume delivery constant
Inspiratory pressure varies
Inspiratory flow constant
Inspiratory time determined by set flow and VT
PCV
Volume delivery varies
Inspiratory pressure constant
Inspiratory flow varies
Inspiratory time set by clinician
Notes: 23
The delivered volume is constant in volume ventilation; in pressure ventilation, volume varies with changes in resistance and compliance. In volume ventilation, inspiratory pressure varies with changes in compliance and resistance; with pressure ventilation, the inspiratory pressure is set and remains constant. Inspiratory flow is constant in volume ventilation but varies in pressure ventilation. In volume ventilation, inspiratory time is determined by the set flow and tidal volume; in pressure ventilation the inspiratory time is set by the clinician.
Let’s move on to our discussion of Pressure Control Ventilation.

## Slide 17: Ventilator Modes
Pressure Regulated Volume Control (PRVC)
  Set parameters: Vt, I:E ratio, frequency, PEEP, FIO2
  Pressure-preset mode with set TI (time cycled)
  Ventilator continually adjusts Pinsp to deliver desired Vt
  ? Best features of VAC (guarantees Vt) and PCV (inspiratory flow rate determined by patient, improving comfort)
  Most common mode used at SCMC

## Slide 18: Ventilator Modes
Pressure-Support Ventilation (PSV)
  Set parameters: Inspiratory pressure, PEEP, FIO2
  Patient must trigger all breaths
  Ventilation is determined by Pinsp, patient determined frequency, patient effort, respiratory mechanics (Crs, R)
  Cycling determined by flow – Pinsp delivered until flow decreasing to 25% of initial flow (flow threshold for cycling can be adjusted)
  Work of breathing may be increased in patients with obstruction
  Often used to “exercise” patients and as a weaning mode

## Slide 19: PSV

## Slide 20: PSV in Obstruction
Expiratory time depends on R,C
Ti varies widely
Some ventilators allow configuration of the threshold for switching off PSV

## Slide 21: Other Modes - Not Discussed
CPAP
SIMV
APRV
Volume or Pressure IRV
HFOV
PAV

## Slide 22: Is One Mode Better than Another?
Studies comparing VAC and PCV have not shown differences in mortality, oxygenation or work of breathing
PCV may improve patient-ventilator synchrony

## Slide 23: Patient/Ventilator Synchrony Volume ventilation delivering a preset flow and volume
1
2
3
4
5
6
30
Paw
cmH2O
Sec
V
.
Patient Effort
Adequate Flow
-10
*Adequate flow gives a smooth, linear rise to pressure
PIP
Notes: 12
12
12
12
12
12
The Pressure-Time curve depicts not only the shape of the curve but displays any PEEP present, the PIP, as well as any inspiratory trigger effort as seen in the second pressure curve.
This Pressure-Time curve shows the different types of pressure curves associated with different types of flow delivery.
a. the linear rapid rise to pressure produced by a square
flow wave
b. the curved gradual rise to pressure resulting from a
descending flow wave
c. the curved “shark’s fin” pressure wave with a long
inspiratory plateau

## Slide 24: Patient/Ventilator Synchrony The patient is outbreathing the set flow
1
2
3
4
5
6
30
Paw
cmH2O
Sec
Patient Effort
-10
*Insufficient flow results in an erratic rise to pressure with a lower PIP
PIP
Air Starvation
What can be done to improve comfort?
Notes: 12
12
12
12
12
12
The Pressure-Time curve depicts not only the shape of the curve but displays any PEEP present, the PIP, as well as any inspiratory trigger effort as seen in the second pressure curve.
This Pressure-Time curve shows the different types of pressure curves associated with different types of flow delivery.
a. the linear rapid rise to pressure produced by a square
flow wave
b. the curved gradual rise to pressure resulting from a
descending flow wave
c. the curved “shark’s fin” pressure wave with a long
inspiratory plateau

## Slide 25: Patient/Ventilator Synchrony The patient is outbreathing the set flow
1
2
3
4
5
6
30
Paw
cmH2O
Sec
Patient Effort
-10
*Insufficient flow results in an erratic rise to pressure with a lower PIP
PIP
Air Starvation
1. Increase inspiratory flow 2. Change mode to PCV
Notes: 12
12
12
12
12
12
The Pressure-Time curve depicts not only the shape of the curve but displays any PEEP present, the PIP, as well as any inspiratory trigger effort as seen in the second pressure curve.
This Pressure-Time curve shows the different types of pressure curves associated with different types of flow delivery.
a. the linear rapid rise to pressure produced by a square
flow wave
b. the curved gradual rise to pressure resulting from a
descending flow wave
c. the curved “shark’s fin” pressure wave with a long
inspiratory plateau

## Slide 26: Understanding the pressure-time waveform using a ‘square wave’ flow pattern
Ppeak
Pres
RET tube
Rairways
Pres
Pplat
Pres
Diaphragm
Ventilator
Time
Pressure
+

## Slide 27: Now let’s look at some different pressure-time waveforms using a ‘square wave’ flow pattern
Scenario # 1
Ppeak = Flow x Resistance + Volume
Compliance
Time
Flow
Normal values:
Ppeak < 40 cm H2O
Pplat < 30 cm H2O
Pres < 10 cm H2O
Ppeak
Pres
Pplat
Pres
Pressure
This is a normal pressure-time waveform with normal peak pressures ( Ppeak), plateau pressures (Pplat ) and airway resistance pressures (Pres)
Time
‘Square wave’
flow pattern

## Slide 28: Waveform showing increased airways resistance
Scenario # 2
Ppeak = Flow x Resistance + Volume + PEEP
Compliance
Time
Flow
e.g. ET tube
blockage
Ppeak
Pres
Pplat
Pres
Pressure
The increase in the peak airway pressure is driven entirely by an increase in the airways resistance pressure. Note the normal plateau pressure.
Time
Normal
‘Square wave’
flow pattern

## Slide 29: Waveform showing high airway resistance due to high flow rates
Scenario # 3
Ppeak = Flow x Resistance + Volume + PEEP
Compliance
Time
Flow
‘Square wave’
flow pattern
e.g. high
flow rates
Ppeak
Pres
Pplat
Pres
Pressure
The increase in the peak airway pressure is driven entirely by an increase in the airways resistance pressure caused by excessive flow rates. Note the shortened inspiratory time and high flow.
Time
Normal
Normal (low)
flow rate

## Slide 30: Waveform showing decreased lung compliance
Scenario # 4
Ppeak = Flow x Resistance + Volume + PEEP
Compliance
Time
Flow
e.g. ARDS
Ppeak
Pres
Pplat
Pres
Pressure
The increase in the peak airway pressure is driven entirely by the decrease in the lung compliance. Increased airways resistance is often also a part of this scenario.
Time
Normal
‘Square wave’
flow pattern